# Assistive-listening verification

PRE-SESSION OPERATIONAL CHECK  /  CHEADLE + RENNICK

A fifteen-minute, repeatable check for induction-loop service, test receivers and seating coverage before every session. Use this procedure at Cheadle auditorium (England) and Rennick auditorium (Scotland).

USE BEFORE EVERY SESSION • RECORD THE RESULT BEFORE DOORS OPEN

# Purpose and boundary

The event checker confirms that the installed assistive-listening service is operational for the session as configured, that representative seating zones are covered, and that receivers intended for issue work. This is an operational readiness check; it is not installation commissioning, a full IEC 60118-4 survey, formal standards testing, or an installation certificate. A completed routine check does not by itself certify legal or technical compliance.

## Who does what

| Role | Accountability |
| --- | --- |
| Event checker | Carries out the checks, records point-level evidence, controls receiver stock and signs the record. |
| Event lead / duty manager | Receives any Fail or Incomplete result before doors and decides the opening arrangement. |
| AV / facilities / designated technical support | Investigates faults beyond basic user checks; contact: ______________________________ |
| Front-of-house lead | Receives any seating restriction or temporary arrangement and helps attendees obtain assistance. |

## Jurisdiction note

Both sites are in Great Britain and operate in the Equality Act 2010 reasonable-adjustment context. This checklist supports reliable service but does not, on its own, discharge that duty. For building/access records, Cheadle is in England (where Approved Document M may be relevant); Rennick is in Scotland (where the Scottish Building Standards Technical Handbook may be relevant). Local installation and compliance records may therefore differ. Keep this operational workflow common unless a commissioned room baseline says otherwise.

## Start point

Start early enough to finish before the stated door time and leave time for a basic fix: latest start = door time minus 15 minutes (or earlier where the room, meter or stock needs setup). If the room is occupied, the meter is unavailable, or a fault remains at the deadline, use the Incomplete/Fail paths on page 4; never call an unfinished check Pass.

# 1  Prepare the room and kit

Use the same signal path and room conditions intended for the session. Tick simple confirmations; write a reading or exception wherever evidence is required.

# Required kit

☐ Induction-loop field-strength meter / calibrated loop test instrument (asset or serial: __________________).

☐ Known-working assistive-listening receiver for the test, plus compatible headphones or earpiece.

☐ Known-good battery or charged receiver, spare battery, and the receiver inventory list.

☐ Stable test tone or controlled speech/test track routed through the same programme path that will feed the loop.

☐ Blank completed-check record, pen/tablet, room point plan or marked seating sketch, and cleaning wipes / disposable earpiece covers.

# Meter and receiver pre-use checks

| Step | Do this |
| --- | --- |
| Meter condition | Inspect case, lead, sensor and display for obvious damage; check adequate battery power. |
| Correct mode | Switch to the loop-testing / magnetic field-strength mode and the range used by the room plan. Do not rely on a generic audio-meter mode. |
| Calibration | Read the in-date calibration label or record. If missing or expired, do not treat readings as a normal Pass: record Incomplete and escalate for an approved calibrated instrument. |
| Test receiver | Check casing, connector and controls; fit a known-good battery/charge; select the loop/telecoil mode required by that receiver. |
| Safe listening | Before wearing headphones/earpiece, turn receiver volume to its lowest safe setting; raise gradually only as needed. |

# Room and signal setup

☐ Power the loop controller and relevant mixer/processor/source as they will be used. Confirm the programme source is actually routed to the loop feed, not only to room loudspeakers.

☐ Put the room in a representative event configuration: normal lighting, displays, projectors, chargers and other electrical equipment that may create interference. Note unusual equipment in the record.

☐ Choose one stable source (test tone or controlled speech) and keep it unchanged while measuring. Mute the source for the background check, then activate it for field and listening checks.

☐ Open the room point plan. If no plan exists, mark at least six repeatable points per room before starting (see page 3).

Safety boundary: temporary staff may use authorised user controls and visible status indicators only. Do not open equipment, bypass safety controls, move fixed wiring, or retune system gain. Escalate instead.

# 2  Measure coverage at repeatable points

A centre-of-room walk-through is not enough in these steel-framed rooms. A satisfactory centre reading alone cannot pass either auditorium. Test every required point independently and record the point ID used below.

# Point plan (use for both rooms)

| Point IDs | Required location |
| --- | --- |
| F1 / F2 | Front audience area: centre and one side/perimeter position. |
| R1 / R2 | Rear audience area: centre and one side/perimeter position. |
| L1 / L2 | Left and right lateral/perimeter seating (use the nearest safe fixed-seat positions). |
| A1 | Accessible seating position(s) intended for attendees using the service. Add A2 if there is more than one distinct accessible zone. |
| S1 / S2 | Known weak, steel-affected, floor-level-change or previously reported risk areas. Add room-specific IDs as needed. |

Minimum: six distributed locations per room, including front, rear, both sides/perimeter, accessible seating and known risk areas. If the local marked plan has more zones, test all of them. Point IDs in the plan must be the same IDs in the record (page 6).

# Ordered measurement

1. With the programme source OFF or muted, place the meter at the first point and check ambient magnetic background noise. Record the value/band. Acceptable standard-practice criterion: no worse than approximately −32 dB relative to 400 mA/m, or the documented commissioned room limit: __________________.

2. Turn on the stable test source and confirm it is feeding the loop controller. Room loudspeaker audio alone is not proof of loop feed.

3. Hold/orient the meter exactly as marked by the instrument (normally the sensor in the stated vertical plane). Use seated ear height, approximately 1.2 m above the floor, at every fixed-seating point; keep height and orientation consistent.

4. Allow the display to stabilise over the defined programme interval (about 10 seconds of steady source) before recording the numerical reading or meter band.

5. Operational field target: nominal 0 dB relative to 400 mA/m, commonly within ±3 dB at each required point, or the documented commissioned baseline/tolerance: __________________. Do not average points to hide a failed seating zone.

6. If a reading is unexpected or borderline, check position and orientation, repeat once, and retain both the initial and repeat readings. A low or unusable zone is identified by location and treated as a failed zone, not absorbed into a room-wide pass.

# Listening confirmation

At more than one zone, including A1 and one known-risk point, use the known-working receiver with compatible headphones/earpiece. Start at low volume. Listen to representative programme speech: it must be intelligible, without hum, buzz, distortion, dropouts or intrusive noise. Record the zones and result. Return receiver volume to a safe starting level after testing.

# 3  Check issued receivers and decide the result

# Receiver inventory control

☐ Count required working receivers: required ______  /  working ______  /  accessories ______. Compare with the event requirement or local stock expectation.

☐ For every receiver intended for issue (or the approved inventory sample rule: __________________), check battery/power, audible output and basic controls. Mark each ID on the record or attached inventory list.

☐ Label and segregate any failed receiver immediately (bag/tray marked NOT FOR ISSUE). Do not return it to working stock until repaired and rechecked.

☐ Clean shared receivers, headphones and earpieces between users; use disposable covers where supplied. Provide discreet attendee instructions, including how to obtain help: FOH contact __________________.

# Result logic

| Outcome | Use when |
| --- | --- |
| PASS | Every required coverage point meets its criterion; background and listening checks pass; every required receiver/accessory is checked and ready; no mandatory item is blank or incomplete. |
| FAIL | A required test was completed and did not meet its criterion (for example, a named low zone, failed receiver or source-chain fault). Record the location/ID and reason. |
| INCOMPLETE / NOT COMPLETED | Required evidence could not be obtained, including an occupied/inaccessible room, missing/damaged/out-of-calibration meter, or unresolved check before doors. Record the reason. |

An Incomplete check must never be recorded or communicated as Pass. Checkboxes are for objective confirmations only; write numerical readings, defect descriptions, retest results and exception notes wherever requested.

# Basic fault check (bounded)

☐ Confirm the test source is playing and routed to the loop; do not assume the loop has failed because the room loudspeaker is audible.

☐ Try the known-good test receiver and battery to distinguish receiver trouble from a room-system trouble.

☐ Observe visible loop-controller/rack fault indicators and record what you see; do not adjust internal controls.

☐ After any authorised basic corrective action, repeat the affected point or receiver check and all relevant overall checks before changing the result to Pass. Keep the original failed reading/result.

# When the deadline is reached

If a required zone, receiver stock or system function has failed and the fault remains at door time, tell the event lead/duty manager before doors, record the status and hand over the opening decision. Escalate to AV/facilities using the editable contact on page 1 and raise maintenance ticket/reference: __________________. If the normal service is unavailable, any substitute receiver, alternative assistive-listening method or restricted seating area must itself be verified before being described as usable. Tell front-of-house staff what restriction/alternative was communicated and where.

# 4  Completed-check record  |  blank reusable form

Complete one record per session. Attach the room sketch or inventory list if more space is needed. File completed records with __________________ (owner/repository) under __________________ (retention reference).

| Session details | Entry |
| --- | --- |
| Event / session ID | ____________________________________________________________ |
| Event date / scheduled start | ________________________ / ________________________ |
| Auditorium + jurisdiction | ☐ Cheadle (England) ☐ Rennick (Scotland) |
| Checker / role | ____________________________________________________________ |
| Actual check start / completion | ________________________ / ________________________ |

# Equipment and source

| Item | Identifier / result |
| --- | --- |
| Field-strength meter asset/serial | ________________________ Calibration in-date/due: __________________ |
| Meter mode/range | ________________________ Background criterion used: __________________ |
| Test receiver ID / battery-power result | ________________________ ☐ good ☐ replace ☐ quarantine |
| Test source / programme path | ____________________________________________________________ |
| Event configuration exceptions | ____________________________________________________________ |

# Coverage and listening evidence

Record a number or meter band at every point. If retested, keep both readings (initial → repeat). Point Pass/Fail is independent; do not average points.

| Point ID / location | Background | Field reading (initial → repeat) | Listening zone/result | Point P/F | Notes / defect |
| --- | --- | --- | --- | --- | --- |
| F1 front centre | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| F2 front side/perimeter | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| R1 rear centre | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| R2 rear side/perimeter | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| L1 left perimeter | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| L2 right perimeter | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| A1 accessible seating | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| S1 steel/weak risk | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| S2 other room-specific risk | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |
| Extra: __________ | ________ | ________________ | ________________ | ☐ P ☐ F | ________________ |

# Completed-check record  |  outcome and handoff

# Receiver issue readiness

| Control | Record |
| --- | --- |
| Required working quantity / accessories | Required ______ Working ______ Accessories ______ ☐ meets requirement ☐ short |
| Receiver IDs checked | ____________________________________________________________ |
| Power/battery, output, controls | ☐ all checked ☐ exception (list IDs): __________________________ |
| Cleaning / covers completed | ☐ yes ☐ no Action: __________________________________________ |
| Volume reset to safe starting level | ☐ yes ☐ no |
| Attendee help instructions available | ☐ yes FOH contact/desk: _______________________________________ |

# Overall result (choose one)

| ☐ PASS | All mandatory points, listening checks and receiver-readiness checks passed; no required item incomplete. |
| --- | --- |
| ☐ FAIL | A completed required check failed. Reason/location required below. |
| ☐ INCOMPLETE / NOT COMPLETED | Required evidence unavailable or unfinished. Reason required below; never represent as Pass. |

| Mandatory reason for Fail or Incomplete | Defects found / affected zone or receiver ID |
| --- | --- |
| ________________________________________________________________________________ | ________________________________________________________________________________ ________________________________________________________________________________ |
| Immediate action taken | Retest result (retain original evidence) |
| ________________________________________________________________________________ | ________________________________________________________________________________ |
| Maintenance ticket / defect reference | AV/facilities escalation contact and time |
| ______________________________ | ________________________________________ |

# Notification and temporary arrangement

| Who was notified | Time | Restriction / alternative communicated |
| --- | --- | --- |
| Event lead / duty manager: __________________ | ________ | ________________________________________________ |
| Front-of-house lead: ________________________ | ________ | ________________________________________________ |
| AV / facilities: ______________________________ | ________ | ________________________________________________ |

Opening decision / handoff notes: __________________________________________________________________________________
____________________________________________________________________________________________________________

Checker signature / initials: __________________________   Event lead acknowledgement: __________________________   Date/time: __________________

## Retention

The event checker files this completed record and any attached point plan/inventory list at the repository above. Unresolved faults remain open against the maintenance reference until AV/facilities closes them.